# AZIMUTH AXIS VERTICALITY MEASUREMENT

## OVERVIEW

This test is to measure and record the verticality of the Azimuth Axis, and to determine the angular deviations from theoretically true azimuth for use as offsets in the controller pointing model.

## AZIMUTH AXIS VERTICALITY MEASUREMENT TEST PROCEDURE

Follow these steps-

- Set the controller for Manual Pendant control. Drive the Turning Head so it is pointing due North to establish a starting measurement position.  Drive the Reflector in elevation to the zenith position.

- Place a precision 10 arcsec level in-line with North/South alignment of Turning Head, and shim it to zero its bubble.

- Drive the Azimuth in the direction which won't encounter a limit for 360deg of rotation and stop to record the bubble reading at the azimuth headings of-                 0, 30, 60, 90, 120, 150, 180, 210, 240,  270, 300, & 330deg. Goal is alignment of ≤30arcsec.

- Last drive the antenna system to 0deg Azimuth heading.  Drive the antenna in Elevation from 88 to 10deg, and record the change in verticality.

| Azimuth heading | 0 | 30 | 60 | 90 | 120 | 150 |
| --- | --- | --- | --- | --- | --- | --- |
| Level measure (division) | 2.60 | 2.76 | 2.80 | 2.80 | 2.85 | 2.90 |

| Azimuth heading | 180 | 210 | 240 | 270 | 300 | 330 |
| --- | --- | --- | --- | --- | --- | --- |
| Level measure (division) | 2.90 | 2.95 | 2.90 | 2.70 | 2.60 | 2.60 |

Axis is Tilt-	__8.2_  milli-degrees

Direction-	__355__  azimuth degrees

Change in verticality from 88 to 10deg-      ___30 arcsec___

SUMMARY OF AZIMUTH VERTICALITY MEASUREMENT RESULTS

| Parameter | Specification | Pass or Fail |
| --- | --- | --- |
| Summary Result of Azimuth Verticality Measurement | NA | Pass |

Test Date:

Start Time:

End Time:

Comments:

By-	  _______________			Date-	__________________

# REFLECTOR SURFACE ACCURACY MEASUREMENT

## OVERVIEW

This test is to measure and record the surface accuracy of the reflector using photogrammetry. Photogrammetry is a 3-dimensional coordinate measuring technique that uses photographs for metrology using the basic principle of triangulation or intersecting lines in space to compute the location of a point. The camera and software have the ability to self-calibrate at the time of the measurement under the environmental conditions that exist by using roll diversity (some shots taken at 90deg to others), and the mathematical means to automatically solve any error terms in the system. Typical accuracies are 25 to 50 microns (.001" to .002") on a 3-meter object, or as in this application shots taken at 3-meters away from the object, and using overlaid photos to piece together the entire image.

## REFLECTOR SURFACE ACCURACY MEASUREMENT TEST PROCEDURE

- Prepare the antenna with the required photogrammetry components-  1380 targets, 28 target codes, 4 datum targets, and a 104in scale bar. Make a setup shot(s) to verify proper camera and flash settings.

NOTE:  Scale Bar used- Brunson 08-CP-300 certified calibration .0002" for 6pc 104" length used.

- Set the controller for Manual Pendant control to be able to position the reflector in azimuth and elevation for the needed photographs to be taken.

- Determine a repeatable scheme whereby photos will be taken at 45deg intervals around the reflector at a desired location approximately 3meters above the rim of the reflector parallel with the focal axis. Using a EWP(elevated work platform) maneuver the photographer to the 8 locations and take approximately 6-8 shots per, overlapping the object in the view finder so that the combined 8 shots can be linked together to represent the entire object.

- Download the photos into the V-Stars Laptop computer and solve the measurements against an IGES file of the theoretical paraboloid specified for the design. Verify accurate scaling, and best-fitting to datum targets. Then use the software function that will best-fit the points to a paraboloid of a fixed (or best-fitted) focal distance. Record the rms of the measurement. Establish a driver file to be used for subsequent measurements.

- Repeat steps 3 & 4 for 15, 50, & 88deg elevation angles. Record in the table below-

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Reflector Surface Accuracy rms | ≤ .015in | Pass |

Results-

| Elevation Angle | 10 | 50 | 88 |
| --- | --- | --- | --- |
| Surface rms | .012 | .012 | .015 |

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Best-Fit Paraboloid Vertex to Nominal Hub Datum Vertex | none | NA |

## Results-

|  | X | Y | Z | Angle* |
| --- | --- | --- | --- | --- |
| 10deg Elevation | .056 | -.090 | -.006 | -.0256˚ |
| 50deg Elevation | .067 | -.048 | -0- | -.0164˚ |
| 88deg Elevation | .067 | .085 | -.002 | -0- |

## *Note-  This is angular droop toward ground due to gravity.

## SUMMARY OF REFLECTOR SURFACE ACCURACY MEASUREMENT

Test Date:

Start Time:

End Time:

Weather Conditions during Test:

Comments:

By-	  _________________		Date-	_______________

# SUB-REFLECTOR AND FEED POSITION AND ALIGNMENT 				MEASUREMENT

## OVERVIEW

This test is to measure and record the positions of the sub-reflector and feed horn using photogrammetry. XYZ locations of targets placed on the sub-reflector and feed horn will be studied to determine position and axial alignment to the best-fit paraboloid reflector shape. The study will be done at 10, 50, and 88deg elevation angles.

## SUB-REFLECTOR/FEED POSITION AND ALIGNMENT MEASUREMENT TEST PROCEDURE

- Use the same photo sessions used in the reflector surface measurement to measure the targets of the feed system.

- Download the photos into the V-Stars Laptop computer and solve the measurements. Create construction lines, planes, and circles from which position, and axial measurements can be determined from.

- Repeat process for the 10, 50, and 88deg elevation angles.

SubReflector  Measurement-

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Sub-reflector position to Best-Fit Paraboloid at nominal 50deg Elevation position | ≤ +/- .08in | Pass |

Results-

|  | X | Y | Z | angle |
| --- | --- | --- | --- | --- |
| Theoretical Position of Sub-reflector Datum Targets | 0 | 0 | 175.45in | -0- |
| Measured Position of Sub-reflector | .013 | -.071 | 175.50 | -0- |

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Sub-reflector position change measurement throughout elevation range of motion | ≤ +/- .120in | P |

Results-

| Elevation Angle | 10 | 50 | 88 |
| --- | --- | --- | --- |
| Sub-reflector Position change from nominal (in) | -.022 | -0- | +.031 |

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Sub-reflector angular alignment change measurement Throughout elevation range of motion | ≤ .05 deg | Pass |

Results-

| Elevation Angle | 10 | 50 | 88 |
| --- | --- | --- | --- |
| Angular change from nominal (deg) | -.013deg | -0- | +.042deg |

Feed System Measurement -

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Feed Horn phase center position to Best-Fit Paraboloid at nominal 50deg Elevation position | ≤ .080in | Pass |

Results-

|  | X | Y | Z | angle |
| --- | --- | --- | --- | --- |
| Theoretical Position of Feed Phase Center | 0 | 0 | 145.21in | NA |
| Measured Position of Feed Phase Center | .013 | -.071 | * | NA |

| ITEM | SPECIFICATION | Pass or Fail |
| --- | --- | --- |
| Feed position change measurement throughout elevation range of motion | ≤ +/- .100in | Pass |

Results-

| Elevation Angle | 10 | 50 | 88 |
| --- | --- | --- | --- |
| Feed Position change from nominal (in) | +.01 | -0- | +.094 |

## SUMMARY OF SUB-REFLECTOR/FEED POSITION AND ALIGNMENT MEASUREMENT

Test Date:

Start Time:

End Time:

Weather Conditions during Test:

Comments:

By-	  __ ________________		Date-	______________